Nanophotonic-Carbohydrate Lab-On-A-Microneedle for Rapid Detection of Human Cystatin C in Finger-Prick Blood

Srinivasu Valagerahally Puttaswamy,† Gennady V. Lubarsky,† Colin Kelsey,† Xushuo Zhang,† Dewar Finlay,† James A. McLaughlin,†,‡ and Nikhil Bhalla∗,†,‡

†Nanotechnology and Integrated Bioengineering Centre (NIBEC), School of Engineering, Ulster University, Jordanstown Shore Road, BT37 0QB, Northern Ireland, United Kingdom

‡Healthcare Technology Hub, Ulster University, Jordanstown Shore Road, BT37 0QB, Northern Ireland, United Kingdom

E-mail: n.bhalla@ulster.ac.uk

Abstract

Miniaturized total analysis systems, for the rapid detection of disease biomarkers, with features including high biomarker sensitivity, selectivity, biocompatibility and disposability, all at low cost are of profound importance in the healthcare sector. Within this frame of reference, we developed a lab-on-carbohydrate-microneedle, biodevice, by integrating localized surface plasmon resonance (LSPR) paper based substrates with biocompatible microneedles of high aspect ratio (>60:1::length:width). These microneedles are completely fabricated with carbohydrate (maltose) and further coated with poly-lactic-co-glycolic acid (PLGA) which together serves the purpose of fluid channels. The porous nature of PLGA, in addition to drawing blood by capillary action, filters outs the whole blood, allowing only the blood plasma to reach the biorecognition layer of the developed biodevice. While the use of maltose provides biocompatibility to the microneedle, the axial compression and transverse load analysis revealed desired mechanical strength of the microneedle, with mechanical failure occurring at 11 N and 9 N respectively for the compressive and transverse load. For proof-of-principle demonstration, the developed biodevice is validated for its operational features by direct detection of Cystatin C in finger-prick blood and up to a concentration of 0.01 µg/ml in buffered conditions using LSPR technique. Furthermore, by changing the biorecognition layer, the use of the developed needle can be extended to other disease biomarkers, and therefore the innovation presented in this work represents a hallmark in the state of the art of lab-on-chip biodevices.

Keywords

POCT, microneedle-lab, LSPR, biosensors, blood-extraction

Advanced point of care testing (POCT) technologies which aim to provide quick treatment, influenced by decisions on patients’ prognosis, must yield analysis within a period of few minutes.1 In clinical diagnostics, most common procedures involve blood sample extraction, followed by analysis of biomolecules of interest through a variety of analytical assays.2,3 Common biochemical assays used for the analysis are enzyme-linked immunosorbent assay (ELISA) for protein, lipid and carbohydrate based molecules4 or the polymerase chain reaction (PCR) for the molecules made up with nucleic acids (DNA/RNA). 5 While ELISA and PCR based technologies are still actively used in the medical setting, a large number of alternative techniques based on the optical6 and electrochemical7 biosensing principles have been developed. In general, attributes such as enhanced sensitivity, selectively and fast response time have been the striking feature of biosensing technology in contrast to the overall performance of conventional ELISA and PCR.8–11 However, it is imperative to integrate blood collection and sampling techniques with advanced biosensors to achieve rapid detection of a target biomolecule, especially at the site of a patient, which is the hallmark of POCT.12 It is in this context that we integrate the localized surface plasmon resonance (LSPR) biosensing technique with a microneedle which painlessly collects the blood sample from a finger and measures the biomarkers within it.

LSPR refers to the optical phenomenon of collective oscillation of valence electrons upon absorption of light within the ultraviolet-visible (UV-Vis) band. This is due to interactions between the incident photons and the conduction band of a noble metal nanostructure.13 These electronic oscillations lead to the generation of an evanescent wave which is sensitive to the local refractive index of the nanostructure.14 This feature acts as the foundation for biosensing applications where either changes in the wavelength of resonance or intensity of the light absorbance can be correlated to the concentration of the binding molecules on the surface of the nanostructure.15 The latest advances in nanotechnology have developed LSPR into a highly sensitive, refractive index specific, and short time response probing method that can quantitate analytes down to even a single molecule level in certain cases16,17

On the other hand, the microneedle technology has been developed for minimally invasive applications, including extraction of blood and interstitial fluid through the skin,18,19 obesity treatment,20 drug delivery,21 insulin injections22 and as neural interfaces.23 Furthermore, microneedles have also been combined with sensors such as enzyme based electrochemical and Surface Enhanced Raman Spectroscopy (SERS) platforms24,25 for the detection of pH,26 lactate27 and alcohol28 in biologically relevant fluids. However, limited work has so far been accomplished for the detection of biomarkers directly from whole blood using microneedles 29,30 Underpinning reasons for the limited work are mainly associated with the challenge of integration of advanced sensing techniques with the microneedle and at the same time ensuring that the microneedle is disposable, biocompatible, pain-free and cost effective.31,32 Besides, there are also critical challenges in the design and fabrication methodology of the microneedles, which include the combined features of a long length, sharp tip and mechanical strength required to reach the blood vessels and extract blood without mechanical failure of the sharp tips.33 Furthermore, incorporating structural features in the microneedle to ensure that an adequate quantity of the blood is extracted for biomarker analysis by the integrated biosensor is still evasive for a lab-on-chip system designer.34 The promising development, nonetheless, is that many of these issues are slowly being addressed by recent advances in the microneedle and biosensing technologies.35–39 However, ensuring that all advantages of the microneedle and biosensing technology are harnessed upon their effective integration in miniaturized total analysis chip formats, remains elusive from both the scientific perspective and technology translation viewpoint.

The developed nanoplasmonic lab-on-microneedle biodevice, see schematic and snapshot in Figure 1a-b, in our work addresses many of the aforementioned challenges. First, the microneedles are fabricated with malt sugar (see schematic in 1c and details in experimental section) which provides the mechanical strength required to penetrate skin and the biocompatibility necessary to interface the biodevice with human tissue.40 The microneedles are further coated with poly-lactic-co-glycolic acid (PLGA) to provide porosity which assists in drawing blood via capillary action.41 In addition, the porous nature of the PLGA provides on-needle filtration of the blood, prior to its advent onto  the LSPR interrogation area. The fabrication of the developed microneedle is also simple, modular, cost effective and readily suitable for mass manufacturing. Essentially the 3D printed moulds, nitrocellulose paper coated with gold nanopartciles-nanoplasmonic paper, the heated solution of malt sugar and the PLGA coating solution are first prepared independently as 4 different modules. The 3D printed microneedle base pillar is then dipped into the sugar solution and slowly drawn out to produce high aspect ratio tips (>60:1:: length: width), see Figure 1c and experimental section for more details. This is followed by sticking the nanoplasmonic paper with a double-sided tape on the 3D printed mould and coating of the microneedle with PLGA solution. All of the materials used in the development of the microneedle are cost effective and easily disposable, with overall mass production, including labor and fabrication related running costs, estimated to be around £0.10 - £0.20 per device for a scale of more than 10,000 units. It should be noted that this estimate does not include the costs for infrastructure such as the cost of 3D printers, the development of injection moulds, or robotic controlled semi-automatic processes for an industrial setting. Nevertheless, it still provides a preliminary cost evaluation concerning technology translation of the developed microneedles.

To demonstrate the proof of concept functioning of the lab-on-microneedle, we detect the biomarker Cystatin C from buffered solutions, spiked human blood samples and directly from finger pricked blood, following informed consent from healthy volunteers. Cystatin C is a protein which is encoded by CST3 gene, primarily used as a biomarker to assess the renal function, for example to diganose chronic kidney disorders (CKD).42 Cystatin C, has also been used to detect the prognostic features of cardiovascular diseases, in particular the onset of heart diseases directly associated with risks of CKD.43 In addition, elevated levels of Cystatin C have also been revealed in Alzheimer’s disease, inflammation and infectious diseases such as SARS-CoV2.44 The normal therapeutic range for the detection of Cystatin C is between 0.6-1 µg/ml, above which the levels are associated with the aberrations of renal function.42,45 In our work we detect Cystatin C successfully at a concentration of 0.01 µg/ml which is more than 10 times lower than the required detection in the clinical settings. The overall time of detection, including sample collection via microneedle and measurement was found to be less than 2 minutes, evidencing the fact that the developed biodevice can potentially be used as a rapid biomarker detection tool right at the site of the patient. Whilst we demonstrate the detection of Cystatin C, this microneedle can also detect other biomarkers by simply changing the biorecognition layers consisting of antibodies or aptamers specific to the target of interest. Therefore, we envisage that the developed biodevice can be used for a number POCT applications, both for diagnosis and prognosis of diseases.

Figure 1: Microneedle schematic, fabrication, and mechanical characteristics: a) schematic depicting microneedles connected to a paper based nanoplasmonic sensing area consisting of gold nanoparticles which upon exposure to white light becomes an active area for the detection of analytes via localized surface plasmon resonance (LSPR); b) shows a snapshot of the microneedle array held in place with forceps and zoomed in picture of a single needle; c) shows the drawing lithography steps i-iv involved in the fabrication of microneedles: i) with a 3D printed mould or supporting substrate and a heated solution of maltose sugar, ii) mould dipped in the sugar and then iii) drawn out of the solution to develop iv) sugar microneedles; d) shows the microneedle performance under the compressive load and e) under the transverse load. Both compressive and transverse strength analysis show that microneedle can withstand load up to 8-12 N.

# Results and discussion

Upon fabrication of the microneedle, compressive and transverse load failure tests to assess axial and shear strengths respectively, of the microneedle were conducted. The compressive test analysis, shown in Figure 1d, revealed a critical load of 11 N which our microneedle can sustain. During this load test, a minute amount of bending in the microneedle was observed between 5-11 N and thereafter needles collapsed completely from the base at 11 N. We also performed a transverse test analysis which indicated that the needle could sustain a maximum load of 9 N, shown in Figure 1e, before it breaks. These force values are more than 10 times higher than the force required to puncture the skin46 indicating that the developed needle has appropriate strength required for the use in the extraction of bodily fluids such as blood. This mechanical strength is attributed to the physical properties of the materials (maltose and PLGA) and the fabrication methodology used to develop the microneedle with desired dimensions. Further details are included in the experimental section. In Figure 2a,b morphological characteristics of a single tip of the microneedle array are shown where Figure 2b is a zoomed-in snap of the image shown in Figure 2a. These scanning electron microscopy (SEM) images clearly depict the porous nature of the microneedle surface due to the PLGA coating. The PLGA coating was observed to be of uniform thickness of 1 µm, from the base to the tip of the microneedle. This porous PLGA coating minimizes the requirement of a stored vacuum system to draw the blood after penetrating through the skin and blood flow occurs due to capillary flow through the interconnected micro structures. In addition, the pore size varies between 0.2-1 µm, with even distribution of pores in the coated maltose surface, allowing plasma filtration, while blocking red blood cells (which are usually about 7 µm in diameter and 2 µm in thickness47) from flowing on to the sensing area, which enhances the specificity of the device. The water absorption test revealed its rapid wetting ability in a short time span of 30 seconds, see more details in the supplementary material including the stability test in liquid environment for a single microneedle tip. The drawing volume of fluid was found to be within 50 µl, for both water and blood, which saturates the maltose/PLGA channels and the nitrocellulose sensing area of the whole biodevice. No drawing of fluid via capillary action was observed beyond the volume of 50 µl. Nevertheless, this volume is within the range required for analyses by direct finger-prick based biodevices such as glucose sensors, mostly extracted by stainless steel microneedles.22

Figure 2c, d shows the distribution of gold nanoparticles (AuNPs) on the nitrocellulose sensing surface of the biodevice.   The SEM in Figure 2c shows a uniform deposition       of nanoparticles on the interwoven structure of the cellulose fibers which constitute the nitrocellulose paper. This interwoven morphology of the paper also presents a large surface area for the attachment of AuNPs via the adsorption process. From Figure 2d it can be observed that the AuNPs deposited on the nitrocellulose paper substrate without forming any aggregates.  The average size of the AuNPs deposited on the paper to form the LSPR active surface was found to be 22.11 ± 5.99 nm, see Figure 2e. The wavelength of LSPR of this developed nanoplasmonic paper was found to be around 512 nm which varies around 1.5% from sample to sample (the worst case lower and upper limit: 506 nm and 520 nm), see Figure 2f including the inset with statistics presented by the violin type graph. These variations are attributed to the minute differences in AuNPs size as observed from Figure 2e and the nature of the adsorption process which leads to a stochastic distribution of AuNPs on the paper substrate. Note that this variability is a hallmark of depositing functionalized nanoparticles on paper-based substrates via direct adsorption process. Direct nanoparticle adsorption by drop casting ensures that the antibody functionalized nanoparticle retains its antigen capturing functionality. It is well-established that the LSPR is dependent on the size and spacing of the nanoparticle distribution on a given surface.13 Ensuring highly reproducible LSPR substrates with a deterministic distribution of nanoparticles using protocols of chemical adsorption remains challenging.

However, these variations in the LSPR signal can be addressed as signal baseline corrections during the acquisition of the LSPR signal where only the relative changes in the wavelength of the LSPR upon biomolecule binding are acquired, instead of correlating the absolute wavelength to the concentration of the binding molecule.16 The sensitivity of the developed LSPR substrate as 133.7 nm  2.9 nm/RIU where  error is the standard deviation, see more details in the supplementary material, section Sensitivity analysis. This standard deviation shows 2.16% deviation in the sensitivity as long as baseline correction is performed. We believe that the 2.16 % deviation is reliable for biosensing applications. Following the morphological, structural characterization of the microneedle and an optical assessment of the LSPR sensing area, bioassay for the detection of Cystatin C was performed to demonstrate the utility of the microneedle for rapid biomarker detection. The detection of Cystatin C was successfully achieved in the buffer solutions, Figure 3, and in whole blood

Figure 2: Microneedle morphology and nanoparticle response: a-b) reveals the scanning electron microscopy image a single microneedle where b) is a zoomed-in image of a) showing porous morphology of the poly-lactic co-glycolic acid (PLGA) coated maltose microneedle; c-d) shows the reaction area of the microneedle with gold nanopartciles distributed on the paper based substrate where d is the zoomed in image of c); e) reveals that the average size of the nanoparticles on the sensor surface is around 22.11 nm a standard divation of 5.11 nm among the developed microneedles; and f) shows the LSPR peak at 512.50 nm from the needle. The inset inside the graph is a violin plot showing the average value of LSPR peak from n=18 samples.

samples as shown in Figure 4. It should be noted that the detection is determined by the presence of the biomolecule within a distance of 10 nm, as the evanescent field of the LSPR has a short decay length. Henceforth, this implies that the paper based nanoplasmonic substrate can detect changes in the refractive index within 10 nm from the surface of the AuNPs. As the size of the antibody and the Cystatin C complex is less than 10 nm, true detection of the antigen, Cystatin C, is achieved by the microneedle device. Furthermore, any salt reminiscent from the buffer (above 10 nm away  from the surface of the AuNP)  is not detected by the LSPR signals, which constitutes the bulk refractive index of the nanoplasmonic paper. Figure 3a shows a typical LSPR curve obtained in our measurements corresponding to the surface of AuNPs with no analytes, Cystatin-C antibody and with antibody-Cystatin C (0.1 µg/ml) complex. LSPR response to a range of concentration from 0.001-3 µg/ml is further shown in Figure 3b, see more details in the supplementary material.

Red shifts in the wavelength of the LSPR upon specific attachment of Cyctain-C on AuNPs are observed, with the sensor surface saturating with values between 35-40 nm of wavelength  shifts at concentrations above 2 µg/ml.  These red shifts in the wavelength  of LSPR are due to the increase in the refractive index primarily attributed to the addition of biomolecule mass on the surface of AuNPs. In Figure 3c, we isolate lower concentrations of Cystatin C to evaluate the limit of detection of the microneedle. The LOD of the microneedle was calculated to be 0.1 µg/ml, see more detailed calculation in supplementary material. This value of LOD is higher than the 1 µg/ml, required to be detected in the therapeutic applications, reflected in several metabolic aberrations in response to a disease. Figure 3c also shows the effect of buffer solution, PBS, on the LSPR response of the microneedle. Surprisingly, a blue shift in the LSPR response of the microneedle was observed when exposed to the PBS (phosphate buffered saline). We attribute this shift to the charge effect on LSPR, rather than the addition of mass on the surface, whereby the charges of the buffer enhance the frequency of oscillation of electrons (as plasmon frequency is proportional to the number of free electrons on the metal surface) leading to a decrease in the wavelength of the LSPR, hence blue shifts.

This charge related effect has also been observed in our previous work on nanoplasmonic biosensors.48,49 We also evaluate other features of the sensor above the LOD, such as the dose response analysis by adding increasing concentrations of Cystatin C from 0.1-3 µg/ml on the same sensor surface to evaluate saturation of the LSPR sensor in Figure 3d and the dynamic range of the sensor, Figure 3e. In Figure 3d, where the concentrations are presented in log scale reconfirmed that the saturation of the LSPR paper substrate occurred at a concentration of 2.5 µg/ml. In addition, the dynamic range of the detection, plotted on a linear scale, fitted with a linear regression line (R2= 0.98) in Figure 3e, was found to exist between 0.1-2 µg/ml.  The specific binding of Cystatin C on  the LSPR active area of the microneedle was further validated by two control experiments. In the first experiment the microneedle was exposed to a high concentration (3 µg/ml) of N-terminal pro B-type natriuretic peptide (NT-proBNP) and in the second experiment Bovine serum albumin (BSA) at a concentration of 3 µg/ml was used to test the microneedle. In both experiments, we observed blue shift compared to the similar concentrations of Cystatin C, see Figure 3f, was noticed, a trend similar to the observation from the effect of PBS buffer. It should be noted that the non-specific attached protein molecules cannot completely be ruled out here as both BSA and NT-proBNP may have adsorbed on the nitrocellulose fibers present in the matrix. However, these proteins do not attach directly to the AuNPs (due to the presence of Cystatin C antibody on its surface) and are only present in the proximity of the AuNPs, for instance in the empty space on the cellulose fiber between two AuNPs. Therefore, their non-specific attachment leads to a change in the charge around the AuNPs rather than a direct mass addition which resulted in blue shifts in the LSPR signal. This suggests that the developed microneedles are highly specific to the target of interest, for example, Cystatin C in our proof of principle detection.

We next extended the use of the microneedle in whole blood. Healthy volunteers were invited for the direct use of a needle in the finger-prick blood. The volunteers reported no pain relative to their experience from conventional syringe needles. The process by which the blood extraction takes place is shown in Figure 4a, b.

Figure 3: LSPR detection of Cystatin C in buffer solution. a) LSPR peaks in wavelength vs. absorbance plot showing the detection of Cystatin C at 0.1 µg/ml; b) shows the change in LSPR wavelength of the micro needle sensing area upon varying concentration of Cystatin C from 0.0001-3 µg/ml; c) highlights the variation of LSPR signal at lower concentrations which are detected using the microneedle. d) shows the variation of the LSPR sensor response on a log scale fitted with a logistic curve to reveal the dose response of the microneedle. Note that lower concentrations i.e. <0.1 µg/ml are not shown in the dose response analysis as it is clear from c) that the lower concentrations cannot be easily distinguished from each other; e) shows the dynamic range of the microneedle on a linear scale, fitted with a linear regression line with R2= 0.98; f) shows LSPR wavelength shifts for control experiments performed with high concentration (3 µg/ml) of NT-proBNP, BSA and PBS against the detection of Cystatin C at a concentration of 1 µg/ml and 3 µg/ml.

Whilst Figure 4a shows the real image of the microneedle in action on a volunteers’ finger, It should be noted that the length of the tips of our microneedle (around 1800 µm) is large enough to penetrate at a depth where blood vessels are accessible (around 1000 µm).50 Moreover if the depth of penetration is beyond 1000 µm, it will stimulate nerve endings, resulting in increased pain.51 Figure 4b shows sequential steps in blood extraction by the microneedle. Upon skin puncture and insertion of needle tips under the skin, the blood is extracted via capillary action. As  previously mentioned, the blood is filtered as it passes through the microneedle towards the nanoplasmonic paper. The blood plasma further moves along the nanoplasmonic paper by capillary forces, wetting the complete LSPR sensor area of the microneedle. This fluid movement can be traced with the naked eye by noticing a colour change in the paper (corresponding to dry/wet conditions) until complete saturation of the surface is achieved. At this point LSPR measurements are conducted and wavelength shifts of Cystatin C detection are reported in Figure 4 c, d.

Figure 4: LSPR detection of Cystatin C in a finger prick blood. a) shows the microneedle being trialled on a human finger; b) shows the steps of how i) the microneedle punctures the skin, ii-iv) draws out the blood from the skin into the needle; c) shows the wavelength shifts upon LSPR detection of Cystatin C directly from the finger (whole blood) and from the whole blood which was collected, then spiked with Cystatin C with a concentration of 1 µg/ml and 3 µg/ml respectively; d) reveals a comparison on microneedle detection in PBS buffer and whole blood. The mean differences in concentrations in buffer and blood show that both concentrations, while easily distinguishable from each other, are also comparable in different fluids (buffer and whole blood). Note: error bars in c and d correspond to replications  of the experiments where n ≥ 3.

Figure 4c shows the wavelength shifts upon LSPR detection of Cystatin C directly from the finger (whole blood) and from the whole blood which was collected, and then later spiked with Cystatin C with a concentration of 1 µg/ml and 3 µg/ml respectively. Wavelength shifts ranging from 27-60 nm were recorded for all of the measurements conducted in blood. We also correlate the LSPR wavelength shifts (27-38 nm) obtained from the un-spiked whole blood with the LSPR wavelength shifts observed with the PBS buffer to concentration of Cystatin C. From Figure 3, the corresponding 23-38 nm of wavelength shifts indicate a concentration of 1-2 µg/ml. It should be noted that we show this comparison only for the purpose of the sensor analysis, and we cannot confirm that the whole blood had a concentration of 1-2 µg/ml.  This is due to the presence of several admixtures in the blood as compared to PBS, the effect of which is not completely isolated in our work. Further, the concentration of filtered blood that reaches the sensor surface will not entirely be homogenous as the nitrocellulose matrix will undergo surface wetting fluid via capillary action. However, the use of such simple paper based analysis is common in lateral flow assay strips where the recognition of the sample analyte results in an appropriate response as also indicated by the experiments in our work.  Detailed clinical evaluation of the microneedle, which is outside the  scope of this work, is required to make more factual assertions on the concentrations. Nevertheless, our experiments form a solid scientific basis for the use of this microneedle to detect biomarkers in whole blood. We also compared mean differences in the LSPR signals from 1 µg/ml and 3 µg/ml concentrations in buffer and blood in Figure 4d. The mean differences in concentrations in the buffer and blood show that both concentrations, while easily distinguishable from each other in the same fluid, are also comparable in different fluids (buffer and whole blood), indicating that our separate calibration (dose-response) may be required for the use of the microneedle in different buffers. However, it still suggests that the microneedle serves as a robust platform for the rapid detection of biomarkers in whole blood.  In addition, we also carried out in-laboratory based stability analysis of the developed microneedle, see figure 5 for details. Essentially, LSPR response was measured from the functionalized microneedles after 1 month of refrigerated storage (approx. 4 oC). No visual degradation of the functionalized nitrocellulose paper was observed, as also indicated by the LSPR measurements in figure 5.

# Figure 5: Stability of the LSPR measurement: variation in the absorption of LSPR wavelength of the nanoplasmonic paper substrate after 1 month and stability of the sensor upon binding of the 1 µg/ml and 3 µg/ml of Cystatin C on the functionalized sensor surface after 2 hours in whole blood.

# The microneedle was deemed suitable for sensing applications and therefore we detected Cystatin C in whole blood spiked with concentrations of 1 µg/ml and 3 µg/ml. After acquiring initial measurement, post 15 min incubation, the microneedle was stored in humidifying chamber for 2 hours (2 h) and response of the microneedle was measured again to check the stability of the measurement. We observe less than 1% shifts in the LSPR characteristics of the microneedle suggesting that the microneedle can be used for rapid detection of Cystatin C without waiting for a long duration for the sensor to stabilize. Although our results suggest good stability, assertions on routine usage of our microneedle in clinical settings must made with caution, until independent tests outside the laboratory setting are performed in near future.

# Conclusions

Even though this microneedle can rapidly detect biomarkers from the blood, subsequent efforts are expected in the field of microneedle based lab on chip systems to translate  the technology for use in routine biomarker detection. For instance, in our work we only demonstrate the use of microneedles as a proof of principle tool for quick confirmation of the disease biomarker and associating prognostic features to this would be impetuous to claim given the current state of the art. Some further topics for investigation in line with our work would include an independent assessment of microneedle in a clinically relevant environment while comparing the performance of the microneedle in parallel with existing tools available to doctors for on-site quick biomarker detection. From a sensor development viewpoint, the development of advanced LSPR materials and their flexibility for integration with on chip biodevices will also be crucial in leading biocompatible microneedles closer to commercialization. In addition, recent works have demonstrated the detection of Cystatin C with electrochemical sensors down to femtograms per milliliter52,53. While these measured concentrations are quite low compared with our work and with what is required in clinical settings, the translation of such devices into rapid detection tools which enable EFA-‘Extraction’-‘Filtration’-‘Analysis’, such as that offered by the microneedle in our work, remains non-trivial. Nonetheless, with advancement in LSPR materials in recent times, especially the sensitivity, we believe that the proof of principle work presented here will set a benchmark for next generation of EFA devices. With some of these improvements, we believe that the microneedles demonstrated in our work could provide an alternative route for biomarker detection, as well as represent an important innovation in the field of biodevices.

# Experimental Section

Materials

All buffer solutions were prepared using de-ionized (DI) water, 18.2 MΩ and all chemicals were used as purchased, unless specified. The following chemicals were used: Maltose monohydrate (BioXtra, ≥99%, molecular weight 360.31), poly-lactic-co-glycolic acid (PLGA, lactide:glycolide 65:35, molecular weight 40000-75000), dichloromethane (DCM, anhydrous, ≥99.8%, contains 40-150 ppm amylene as stabilize), Phosphate buffered saline (PBS, Sigma- Aldrich, UK), filter paper (Whatman, grade1), Cystatin C antibody ab133495 (Abcam, Rabbit IgG monoclonal antibody, 100 µl,  0.276 mg/ml) and Cystatin C antigen ab87481(Recombinant Human Cystatin C protein, Abcam, 10 µg, 0.185 mg/ml), Tetrachlorouronic acid, trisodium citrate were purchased from Sigma-Aldrich, UK.

## Fabrication of plasmonic substrate holder with 3D printed base pillar

The plasmonic substrate holder which also consists of a base pillar for the growth of microneedles, was designed using solidworks computer aided design (CAD) package, see more details in the supplementary material. The design was converted into STL format and later printed using a commercial 3D printer (3D Systems, ProJet 1200). This 3D printer is a stereo lithographic system that delivers fast, high resolution output whilst it has the capabilities of mass production at a small scale (<10,000 units a day) if required for our microneedles. The material used for the microneedle substrate holder with the base pillar is a resin known as VisiJet FTX Green material. This resin is suitable for the fabrication of microscale mechanical components with high the high dimensional precision. The resulting parts were visually assessed using an optical microscope to reveal structural details including smooth patterns on the surface of the base pillar.

## Fabrication of maltose microneedle

The fabrication of maltose microneedles adhering to the 3D printed base pillar is summarized in Figure 1c. Powdered maltose weighing 1 gram was dissolved in 1 ml of DI water and heated on a hotplate at 130o C for 40 min, in a glass container. During this period water was vaporized to form sticky liquid maltose. The temperature was gradually increased from 130o C to 150o C and the 3D printed substrate with base pillar, mounted on a drawing system, was brought close to the glass container with sticky maltose as represented in the Figure 1c (i). The temperature of the liquid maltose was further increased to 170o C while the base pillar tip was still in close contact with the maltose (Figure 1c (ii)), followed by slowly drawing out the base pillar away from the liquid maltose (Figure 1c (iii)). The drawing of the substrate was performed by a customized laboratory-based drawing system. This drawing of the liquid maltose was performed in 3 steps. The substrate was pulled in an upwards direction, first at a low drawing speed of 250 µm/s for one minute, , and then at an increased drawing speed of 350 µm/s for 3 minutes, away from the liquid maltose. This was followed by a much quicker drawing procedure performed at a speed of 600 µm/s. Since the maltose liquid is less viscous at higher temperatures, the connection between the 3D printed microneedle base pillar and surface of the liquid maltose shrinks when drawn at a speed of 600 µm/s. Finally, it forms a neck and then breaks to form a sharp conical tip as illustrated in (Figure 1iv).

## Porous PLGA coating on maltose microneedle array

The PLGA coating solution was prepared using the following procedure as reported by Omar et al..41  At  first,  2% (w/w) PLGA in the crystal form was  dissolved in DCM. Further,   1% of PBS was added to the PLGA/DCM as a porogen and the PLGA/DCM/PBS was homogenized at 20000 rpm for 3 minutes to prepare a coating solution. Just before dip coating of the microneedle with PLGA, the microneedle was placed in a plasma chamber and subjected to oxygen plasma treatment for 15 minutes to enhance the adhesion of the PLGA layer. The plasma treatment conditions used for treating maltose microneedles were as follows: oxygen flow rate of 70 sccm, oxygen pressure 0.2 mTorr, plasma power 200 W and a duration of treatment of 15 minutes. The plasma treated maltose microneedles were dipped into the coating solution for 5 minutes for uniform coating with porous PLGA polymer layer. The PLGA coated maltose microneedle was then immersed in EtOH-NaOH (ethanol sodium hydroxide) solution for 2 minutes to enhance porosity on the microneedle surface. 54

## Scanning electron microscopy analysis of microneedle array

The morphology of the maltose microneedle coated with PLGA was characterized with a Hitachi SU5000 field emission scanning electron microscopy (FESEM) in high-vacuum mode at 5 kV accelerating voltage. Prior to maltose microneedle mounting on a circular stub, sputter- coating with palladium film of thickness 20 nm using Emitech K500X sputtering system to avoid electron charging during imaging was performed.

## Mechanical characterization of the microneedle array

The maltose microneedle array was tested for axial compression and transverse failure forces using a micro-mechanical test machine (InstronH Model 5969; Instron, Norwood, MA). During the axial compression test, the maltose microneedle array was mounted on upper station and pressed against a flat block at 50 µm/s during each test. While conducting a transverse failure test, the maltose microneedle array was placed on a specimen holder, with the microneedle placed perpendicular to the direction of the movement of the upper station with a test speed of 50 µm/s. The force applied and the corresponding displacement was noted to plot a force versus displacement curve to determine the failure load during both tests.

## Plasmonic paper microneedle patch preparation

## Gold nanoparticles (AuNPs) were synthesized by citrate reduction of HAuCl4. In summary, a gold salt solution was prepared by mixing 0.16 g of Tetrachloroauric acid salt with 1600 ml of distilled water. In particular, we would like to suggest that a plastic spatula should be used to handle this gold salt carefully, as stainless steel has a chemical reaction with this salt leading to unwanted impurities in the gold nanoparticle solution. The use of a plastic spatula is also suggested in the material safety and data sheet (MSDS), which would be easily overlooked since no reason is cited. A sodium citrate solution was also prepared by mixing 1g of sodium citrate to 100 ml of DI water. 400 ml of gold salt solution was poured into a thoroughly cleaned round bottom flask, kept on the heating mantle, and gradually heated to a temperature of 100o C. It was allowed to boil for 15 minutes, while stirring with a magnetic flea. During this stage, 9 ml of sodium citrate solution was added to the round bottom flask, which results in a colour change from clear to dark blue/black to wine red in a time span of 30 to 60 seconds. The heating mantle was brought to 25o C, after a further 60 minutes of boiling under reflux. The mixture was allowed to cool to room temperature, 25o C, before labeling and appropriate storage for subsequent analysis/functionalization. AuNPs were directly functionalized with 10 µg/ml of the Cystatin C antibody by means of surface adsorption. Nitrocellulose paper was then dipped inside the AuNPs solution for 1 hour for uniform coating of paper with functionalized gold nanoparticles. The nitrocellulose paper is white in color and consists of a mesh of cellulose fibers which trap nanoparticles efficiently and upon nanoparticle trapping the paper turns pinkish red in colour to indicate that the plasmonic paper substrate is ready for use as a biosensor. With the help of a double-sided tape the plasmonic paper was attached to the 3D printed holder with PLGA coated maltose microneedle in direct contact with the plasmonic paper.

## LSPR signal acquisition and data analysis

The LSPR signal was acquired using a homemade bench top absorbance setup,also used in our previous works.48,55 This setup consists of components purchased from ocean optics: spectrometer FLAME-T-XR1-ES, reflection probe QR400-7-SR-BX, HL-2000 tungsten halogen lamp, RTL-T stage and Ocean View software. Prior to LSPR measurements dark and reference spectra for background noise cancellation were measured using blank nitrocellulose paper as a reference. All generated data was analyzed and plotted using the in- built functionality of the Graphpad prism 8 software.

## Bioassay protocol

Cystatin C antigen was diluted with PBS in different concentrations according to our range (0.0001 µg/ml to 3 µg/ml). The concentration of the stock solution was 185 µg/ml. To obtain the concentration of 3 µg/ml of Cystatin C, 2 µl stock solution was diluted with 30.83	µl of PBS. Subsequently, the lower concentrations of Cystatin C antigen such as 2.5, 2, 1.5, 1, 0.5, 0.1, 0.01, 0.001 and 0.0001 µg/ml were obtained by serial dilution of the previous concentration with an appropriate volume of PBS buffer. Two µl of capture antibody stock solution with a concentration of 276 µg/ml was diluted with 55.2 µl of PBS to obtain the required concentration of 10 µg/ml.

It was directly adsorbed onto AuNPs surface by adsorption. Further, the nitrocellulose paper samples were cut to the required size and dipped into the AuNPs solutions for 24 hours. It was taken out and allowed to dry for 30 minutes at ambient temperature. AuNPs functionalized with capture antibody was embedded in the 3D structure, uniformly spread and adsorbed into the paper as shown in supplementary figures at different locations. The van der Waals force causes high adsorption of AuNPs on the paper without the need for any external aid. At first, LSPR reading was taken, with AuNPs on paper without capture antibody. In the next step, LSPR reading was taken with AuNPs on paper functionalized with the capture antibody. Further, 5 µL of Cystatin C with a minimum concentration of 0.0001 µg/ml was placed on the glass slide. PLGA coated porous malt sugar microneedles were brought in contact with the solution to quickly transport it to the sensing interrogation area by capillary action and an LSPR reading was taken. Later on, the LSPR substrate was washed with PBS buffer and allowed to dry. Subsequently, a similar procedure was repeated while taking LSPR readings each time we increased the concentration of Cystatin C.

## Blood extraction procedure

We tested our device on an index finger, which was punctured with the microneedle at the center of the finger pad, following informed consent within the Ulster University health and safety guidelines. The finger was cleaned with an alcohol swab at the pricking site and the alcohol was allowed to dry. The microneedle was gently pushed into the skin until it reaches a depth of 1000 µm to access blood capillaries. The microneedle array consists of 10 microneedles to puncture an adequate number of blood capillaries to obtain a good blood flow. When the microneedle penetrates human skin and punctures the blood vessels, the blood first comes in contact with the porous layer of the maltose microneedle, which by capillary action allows blood to flow onto sensing area of the device.

# Supporting Information

Following supporting information is freely available at the ACS publication website:

Details on microneedle holder design; comprehensive SEM of tip surfaces and distribution of gold nanoparticles on paper substrates; set up for mechanical characterization of the microneedle; LSPR dose response, limit of detection and sensitivity analysis including stability of the microneedle in liquid environment and a video demonstrating the liquid flow  within the microneedle.

# Author contributions

SVP performed all experiments related to sensing, characterization and design of the microneedle. GL and XZ assisted SVP in experiments related to microneedle fabrication and LSPR sample preparation. DF, JM, NB planned and supervised the overall work. NB prepared the initial manuscript draft with input from SVP. SVP, DF, JM, CK and NB revised the final version of the manuscript and all authors approved the submission.

# Acknowledgments

JM and SVP would like to gratefully acknowledge the funding under the Invest Northern Ireland and the European Union’s INTERREG VA Programme, managed by the Special EU Programmes Body (SEUPB) under the Connected Health Innovation Centre (CHIC) competence center. JM and NB would like to acknowledge EPSRC support on this project from the grant EP/TO24437/1 to the Ulster University. The authors would also like to thank Dr. Preetam Sharma at NIBEC, Ulster University for his help in the SEM imaging.

# References

    - Rickard, J. J.; Di-Pietro, V.; Smith, D. J.; Davies, D. J.; Belli, A.; Oppenheimer, P. G. Rapid Optofluidic Detection of Biomarkers for Traumatic Brain Injury via Surface-Enhanced Raman spectroscopy. Nat. Biomed. Eng. 2020, 4, 1–14.

    - Tsuchiya, K.; Nakanishi, N.; Uetsuji, Y.; Nakamachi, E. Development of Blood Extraction System for Health Monitoring System. Biomed. Microdevices 2005, 7, 347–353.

    - Meldgaard, M.; Bollen, P. J. A.; Finsen, B. Non-Invasive Method for Sampling and Extraction of Mouse DNA for PCR. Lab. Anim-UK 2004, 38, 413–417.

    - Lequin, R.M. Enzyme Immunoassay(EIA)/Enzyme-Linked Immunosorbent Assay (ELISA). Clin. Chem. 2005, 51, 553–556.

    - Saiki, R.; Gyllenstein, U.; Erlich, H. Polymerase Chain Reaction. Science 1988, 239, 487.

    - Borisov, S. M.; Wolfbeis, O. S. Optical Biosensors. Chem. Rev. 2008, 108, 423–461.

    - Ronkainen, N. J.; Halsall, H. B.; Heineman, W. R. Electrochemical Biosensors. Chem. Soc. Rev. 2010, 39, 1747–1763.

    - Sreekanth, K. V.; Alapan, Y.; ElKabbash, M.; Ilker, E.; Hinczewski, M.; Gurkan, U. A.; Luca, A. D.; Strangi, G. Extreme Sensitivity Biosensing Platform Based on Hyperbolic Metamaterials. Nat. Mater. 2016, 15, 621–627.

    - Del Río, J. S.; Henry, O. Y.; Jolly, P.; Ingber, D. E. An Antifouling Coating  that Enables Affinity-Based Electrochemical Biosensing in Complex Biological Fluids. Nat. Nanotechnol. 2019, 14, 1143–1149.

    - Lei, Y.; Sun, R.; Zhang, X.; Feng, X.; Jiang, L. Oxygen-Rich Enzyme Biosensor Based on Superhydrophobic Electrode. Adv. Mater. 2016, 28, 1477–1481.

    - Zhao, L.; Li, H.; Meng, J.; Wang, A. C.; Tan, P.; Zou, Y.; Yuan,  Z.; Lu, J.; Pan, C.;  Fan, Y. Reversible Conversion between Schottky and Ohmic Contacts for Highly Sensitive, Multifunctional Biosensors. Adv. Funct. Mater. 2020, 30, 1907999.

    - He, R.; Niu, Y.; Li, Z.; Li, A.; Yang, H.; Xu, F.; Li, F. A Hydrogel Microneedle Patch for Point-of-Care Testing Based on Skin Interstitial Fluid. Adv. Healthc. Mater. 2020, 9, 1901201.

    - Hammond, J. L.; Bhalla, N.; Rafiee, S. D.; Estrela, P. Localized Surface Plasmon Resonance as a Biosensing Platform for Developing Countries. Biosensors 2014, 4, 172–188.

    - Im, H.; Bantz, K. C.; Lee, S. H.; Johnson, T. W.; Haynes, C. L.; Oh, S.-H. Self-Assembled Plasmonic Nanoring Cavity Arrays for SERS and LSPR Biosensing. Adv. Mater. 2013, 25, 2678–2685.

    - Park, Y.; Ryu, B.; Deng, Q.; Pan, B.; Song, Y.; Tian, Y.; Alam, H. B.; Li, Y.; Liang, X.; Kurabayashi, K. An Integrated Plasmo-Photoelectronic Nanostructure Biosensor Detects an Infection Biomarker Accompanying Cell Death in Neutrophils. Small 2020, 16, 1905611.

    - Bhalla, N.; Sathish, S.; Galvin, C. J.; Campbell, R. A.; Sinha, A.; Shen, A. Q. Plasma-Assisted Large-Scale Nanoassembly of Metal-Insulator Bioplasmonic Mushrooms. ACS Appl. Mater. Interfaces 2018, 10, 219–226.

    - Lednicky, T.; Bonyár, A. Large Scale Fabrication of Ordered Gold Nanoparticle-Epoxy Surface Nanocomposites and Their Application as Label-Free Plasmonic DNA Biosensors. ACS Appl. Mater. Interfaces 2020, 17, 4804-4814

    - Zhang, X.; Chen, G.; Bian, F.; Cai, L.; Zhao, Y. Encoded Microneedle Arrays for Detection of Skin Interstitial Fluid Biomarkers. Adv. Mater. 2019, 31, 1902825.

    - Chang, H.; Zheng, M.; Yu, X.; Than, A.; Seeni, R. Z.; Kang, R.; Tian, J.; Khanh, D. P.; Liu, L.; Chen, P. A Swellable Microneedle Patch to Rapidly Extract Skin Interstitial Fluid for Timely Metabolic Analysis. Adv. Mater. 2017, 29, 1702243.

    - Zhang, Y.; Liu, Q.; Yu, J.; Yu, S.; Wang, J.; Qiang, L.; Gu, Z. Locally Induced Adipose Tissue Browning by Microneedle Patch for Obesity Treatment. ACS Nano 2017, 11, 9223–9230.

    - Kim, S.; Yang, H.; Eum, J.; Ma, Y.; Lahiji, S. F.; Jung, H. Implantable Powder- Carrying Microneedles for Transdermal Delivery of High-Dose Insulin with Enhanced Activity. Biomaterials 2020, 232, 119733.

    - Yu, J.; Wang, J.; Zhang, Y.; Chen, G.; Mao, W.; Ye, Y.; Kahkoska, A. R.; Buse, J. B.; Langer, R.; Gu, Z. Glucose-Responsive Insulin Patch for the Regulation of Blood Glucose in Mice and Minipigs. Nat. Biomed. Eng. 2020, 4, 499-506.

    - Xiang,  Z.;  Yen,  S.-C.;  Xue,  N.;  Sun,  T.;  Tsang,  W.  M.;   Zhang,  S.;   Liao,  L.-  D.;  Thakor, N. V.;  Lee, C. Ultra-Thin Flexible Polyimide Neural Probe Embedded      in a Dissolvable Maltose-Coated Microneedle. J. Micromech. Microengin. 2014, 24, 065015.

    - Park, J. E.; Yonet-Tanyeri, N.; Ende, E. V.; Henry, A.-I.; White, B. E. P.; Mrksich, M.; Duyne, R. P. V. Plasmonic Microneedle Arrays for In Situ Sensing with Surface-Enhanced Raman Spectroscopy (SERS). Nano Lett. 2019, 19, 6862–6868.

    - Kolluru, C.; Gupta, R.; Jiang, Q.; Williams, M.; Derami, H. G.; Cao, S.; Noel, R. K.; Singamaneni, S.; Prausnitz, M. R. Plasmonic Paper Microneedle Patch for On-Patch Detection of Molecules in Dermal Interstitial Fluid. ACS Sens. 2019, 4, 1569–1576.

    - Mani, G. K.; Miyakoda, K.; Saito, A.; Yasoda, Y.; Kajiwara, K.; Kimura,  M.;  Tsuchiya, K. Microneedle pH Sensor: Direct, Label-Free, Real-Time Detection of Cerebrospinal Fluid and Bladder pH. ACS Appl. Mater. Interfaces 2017, 9, 21651–21659.

    - Bollella, P.; Sharma, S.; Cass, A. E. G.; Antiochia, R. Microneedle-Based Biosensor for Minimally-Invasive Lactate Detection. Biosens. Bioelectron. 2019, 123, 152–159.

    - Mohan, A. V.; Windmiller, J. R.; Mishra, R. K.; Wang, J. Continuous Minimally-Invasive Alcohol Monitoring Using Microneedle Sensor Arrays. Biosens. Bioelectron. 2017, 91, 574–579.

    - Ventrelli, L.; Strambini, L. M.; Barillaro, G. Microneedles for Transdermal Biosensing: Current Picture and Future Direction. Adv. Healthc. Mater. 2015, 4, 2606–2640.

    - Li, C. G.; Joung, H.-A.; Noh, H.; Song, M.-B.; Kim, M.-G.; Jung, H. One-Touch-Activated Blood Multidiagnostic System Using a Minimally Invasive Hollow Microneedle Integrated with a Paper-Based Sensor. Lab Chip 2015, 15, 3286–3292.

    - Miller, P. R.; Narayan, R. J.; Polsky, R. Microneedle-Based Sensors for Medical Diagnosis. J. Mater. Chem. B 2016, 4, 1379–1383.

    - Miller, P. R.; Skoog, S. A.; Edwards, T. L.; Lopez, D. M.; Wheeler, D. R.; Arango, D. C.; Xiao, X.; Brozik, S. M.; Wang, J.; Polsky, R. Multiplexed Microneedle-Based Biosensor Array for Characterization of Metabolic Acidosis. Talanta 2012, 88, 739–742.

    - Dardano, P.; Rea, I.; Stefano, L. D. Microneedles Based Electrochemical Sensors: New Tools for Advanced Biosensing. Curr. Opin. Electrochem. 2019, 17, 121-127.

    - Blicharz, T. M.; Gong, P.; Bunner, B. M.; Chu, L. L.; Leonard, K. M.; Wakefield, J. A.; Williams, R. E.; Dadgar, M.; Tagliabue, C. A.; Khaja, R. E. Microneedle-Based Device for the One-Step Painless Collection of Capillary Blood Samples. Nat. Biomed. Eng.

2018, 2, 151–157.

    - Teymourian, H.; Moonla, C.; Tehrani, F.; Vargas, E.; Aghavali, R.; Barfidokht, A.; Tangkuaram, T.; Mercier, P. P.; Dassau, E.; Wang,  J. Microneedle-Based Detection of Ketone Bodies along with Glucose and Lactate: Toward Real-Time Continuous Interstitial Fluid Monitoring of Diabetic Ketosis and Ketoacidosis. Anal. Chem. 2019, 92, 2291–2300.

    - Prausnitz, M. R.; Gomaa, Y.; Li, W. Microneedle Patch Drug Delivery in the Gut. Nat. Med. 2019, 25, 1471–1472.

    - Gowers, S. A.; Freeman, D. M.; Rawson, T. M.; Rogers, M. L.; Wilson,  R.  C.;  Holmes, A. H.; Cass, A. E.; O’Hare, D. Development of a Minimally Invasive Microneedle-Based Sensor for Continuous Monitoring of β-Lactam Antibiotic Concentrations In Vivo. ACS Sens.2019, 4, 1072–1080.

    - Zhao, L.; Wen, Z.; Jiang, F.; Zheng, Z.; Lu, S. Silk/polyols/GOD Microneedle Based Electrochemical Biosensor for Continuous Glucose Monitoring. RSC Adv. 2020, 10, 6163–6171.

    - Li, C. G.; Lee, K.; Lee, C. Y.; Dangol, M.; Jung, H. A Minimally Invasive Blood- Extraction System: Elastic Self-Recovery Actuator Integrated with an Ultrahigh-Aspect- Ratio Microneedle. Adv. Mater. 2012, 24, 4583–4586.

    - Lee, K.; Lee, C. Y.; Jung, H. Dissolving Microneedles for Transdermal Drug Administration Prepared by Stepwise Controlled Drawing of Maltose. Biomaterials 2011, 32, 3134–3140.

    - Qutachi, O.; Vetsch, J. R.; Gill, D.; Cox, H.; Scurr, D. J.; Hofmann, S.; Müller, R.; Quirk, R. A.; Shakesheff, K. M.; Rahman, C. V. Injectable and Porous PLGA Microspheres that Form Highly Porous Scaffolds at Body Temperature. Acta Biomater. 2014, 10, 5090–5098.

    - Shardlow, A.; McIntyre, N. J.; Fraser, S. D.; Roderick, P.; Raftery, J.; Fluck, R. J.; McIntyre, C. W.; Taal, M. W. The Clinical Utility and Cost Impact of Cystatin C Measurement in the Diagnosis and Management of Chronic Kidney Disease: A Primary Care Cohort Study. PLoS Med. 2017, 14, e1002400

    - Jang, S. Y.; Yang, D. H.; Kim, H. J.; Park, B. E.; Park, Y. J.; Kim, H. N.; Kim, N. K.; Bae, M. H.; Lee, J. H.; Park, H. S. Prognostic Value of Cystatin C-Derived Estimated Glomerular Filtration Rate in Patients with Acute Heart Failure. Cardiorenal Med. 2020, 10, 232–242.

    - Gao, Y.; Li, T.; Han, M.; Li, X.; Wu, D.; Xu, Y.; Zhu, Y.; Liu, Y.; Wang, X.; Wang, L. Diagnostic Utility of Clinical Laboratory Data Determinations for Patients with the Severe COVID-19. J. Med. Virol. 2020, 7, 791-796

    - Filler, G.; Bökenkamp, A.; Hofmann, W.; Bricon, T. L.; Martínez-Brú, C.; Grubb, A. Cystatin C as a Marker of GFR—History, Indications, and Future Research. Clin. Biochem. 2005, 38, 1–8.

    - Jiang, S.; Li, P.; Yu, Y.; Liu, J.; Yang, Z. Experimental Study of Needle-Tissue Interaction Forces: Effect of Needle Geometries, Insertion Methods and Tissue Characteristics. J. Biomech. 2014, 47, 3344–3353.

    - Puttaswamy, S. V.; Fishlock, S. J.; Steele, D.; Shi, Q.; Lee, C.; McLaughlin, J.  Versatile Microfluidic Platform Embedded with Sidewall Three-Dimensional Electrodes for Cell Manipulation. Biomed. Phys. Eng. Express 2019, 5, 055003.

    - Roether, J.; Chu, K.-Y.; Willenbacher, N.; Shen, A. Q.; Bhalla, N. Real-Time Monitoring of DNA Immobilization and Detection of DNA Polymerase Activity by a Microfluidic Nanoplasmonic Platform. Biosens. Bioelectron. 2019, 142, 111528.

    - Bhalla, N.; Sathish, S.; Sinha, A.; Shen, A. Q. Large-Scale Nanophotonic Structures for Long-Term Monitoring of Cell Proliferation. Adv. Biosyst. 2018, 2, 1700258.

    - Braverman, I. M. The Cutaneous Microcirculation. J. Invest. Derm. Symp.  P. 2000, 5, 3–9.

    - Gill, H. S.; Denson, D. D.; Burris, B. A.; Prausnitz, M. R. Effect of  Microneedle Design on Pain in Human Volunteers. Clin. J. Pain 2008, 24, 585–594.

    - Zhao, M.; Bai, L.; Cheng, W.; Duan,X,; We, H.; Ding,S.; Monolayer Rubrene Functionalized Graphene-Based Electrochemiluinsescence Biosensor for Serun Cystation C Detection withImmunorecognition-Induced 3D DNA Machine. Biosens. Bioelectron. 2019, 127, 126-134

    - Hassanain, W.A.; Izake, E.L.; Ayoko, G.A.; Spectroelectrochemical Nanosensor for the Determination of Cystatin C in Human Blood. Anal. Chem. 2018, 90, 10843–10850

    - Bible, E.; Chau, D. Y.; Alexander, M. R.; Price, J.; Shakesheff, K. M.; Modo, M. Attachment of Stem Cells to Scaffold Particles for Intra-Cerebral Transplantation. Nat. Protoc. 2009, 4, 1440.

    - Bhalla, N.; Lee, D.; Sathish, S.; Shen, A. Q. Dual-Mode Refractive Index and Charge Sensing to Investigate Complex Surface Chemistry on Nanostructures. Nanoscale  2017,  9, 547–554.